EksoNR Safety Information

# Environment For Use

EksoNR must be used in appropriate and safe environments:

Solid, dry, level ground with less than a 2% grade

Adequate supervision provided.

Use of an assistive device.

## Indications for EksoNR Use

For customers outside of the US and Canada (CE mark):

The EksoNR is intended for use as a gait training device to improve walking function and independence in patients with a neurological or muscular injury, illness, or weakness. It is intended to help facilitate the restoration or improvement of ambulation for its pilots under the supervision of a trained physical therapist (or equivalent medical professional) and operated by a Trained Spotter.

For customers in the United States and Canada (FDA clearance):

The EksoNR is intended to perform ambulatory functions in rehabilitation institutions under the supervision of a trained physical therapist for the following populations:

Individuals with multiple sclerosis (upper extremity motor function of at least 4/5 in at least one arm)

Individuals with acquired brain injury, including traumatic brain injury and stroke (upper extremity motor function of at least 4/5 in at least one arm).

Individuals with spinal cord injuries at levels T4 to L5 (upper extremity motor function of at least 4/5 in both arms).

Individuals with spinal cord injuries at levels of C7 to T3 (ASIA D with upper extremity motor function of at least 4/5 in both arms).

- The therapist must complete a training program prior to use of the device. The devices are not intended for sports or stair climbing.

WARNING Intended Use

The EksoNR may only be used under the direction of a physical therapist who has been certified by Ekso Bionics, Inc. to operate the EksoNR.

## Patient Requirements

Be screened and cleared by a physician prior to physical therapist evaluation for EksoNR use.

Physical and cognitive ability to use crutches, walker, or cane to participate in initiation and maintaining of weight shift. Use does not need to be independent of clinical support.

Ability to communicate pain and need to cease session, verbally or nonverbally.

Ability to acknowledge communication from the therapist, verbally or nonverbally.

Healthy bone density.

Skeleton does not suffer from any unhealed fractures.

Able to stand using a device such as a standing frame.

Hip width or leg segment lengths are within the range of adjustability.

Weight does not exceed 220 lb (100 kg).

Standing hip width of 18 in. (45,6 cm) or less.

Must have functional use of at least 2 extremities.

Must have near normal range of motion in hips, knees, and ankles.

- Neutral ankle dorsiflexion.

NOTE: The knee may flex up to 12° to obtain a neutral ankle.

- No more than 12° knee flexion contracture.

- No more than 17° hip flexion contracture.

Upper leg length discrepancy must be equal to or less than half an inch (≤ 0.5 in. or 1,3 cm) and lower leg discrepancy equal to or less than three-quarters of an inch (≤ 0.75 in. or 1,9 cm).

## Contraindications for Use

People with the following conditions should not use the device:

Severe concurrent medical conditions: infections, circulatory, heart or lung, pressure sores

Severe spasticity (Modified Ashworth 4)

Unstable spine or unhealed limbs or pelvic fractures

Active heterotopic ossification interfering with lower extremity range of motion

Significant contractures

Psychiatric or cognitive situations that may interfere with proper operation of the device

Cognitive impairments resulting in inability to follow directions

Pregnancy

Poor skin integrity in areas in contact with the device

Decreased standing tolerance due to orthostatic hypotension

Range of motion restrictions that would prevent a patient from achieving a normal, reciprocal gait pattern, or would restrict a patient from completing normal sit-to-stand or stand-to-sit transitions:

- Knee flexion contracture greater than 12°

- Inability to achieve 0° neutral angle dorsiflexion with knee flexion up to 12°

Leg length discrepancy greater than 0.5 inches (1,3 cm) for upper legs or 0.75 inches (1,9 cm) for lower legs

Unresolved deep vein thrombosis

Uncontrolled autonomic dysreflexia

Lower limb prosthesis

## General Warnings

Patients with the following conditions may have additional risks:

Significant osteoporosis that may increase the risk of fractures caused by standing or walking. Patients should be evaluated by their physician to assess severity of osteoporosis and whether they can tolerate walking and standing prior to using the device.

Recent history of uncontrolled orthostasis. Patients’ blood pressure should be closely monitored to minimize risk associated with syncopal episodes.

Hip dysplasia or hip axis abnormalities. Use of the device should be paused and checked for fit if the patient complains of hip discomfort or pain. If the fit is correct and the patient continues to experience discomfort, this device may not be appropriate for use with this patient.

Patients with cranial defect or with cranial bone flap stored in abdomen should be wearing a helmet at all times and should be closely monitored when using the EksoNR.

Impulsive behavior that may lead to unsafe turns, walking, or EksoNR use that may lead to injury to patient and physical therapist. If the patient exhibits cognitive or psychiatric concerns, which can lead to an inability to follow instructions, this device should not be used.

If there are any unsafe changes in the patient’s blood pressure or heart rate, the session should be immediately terminated.

Only use the EksoNR on solid, dry, level ground with less than a 2% grade. Using the device on irregular surfaces could result in loss of balance and possible injury.

The device must be fitted over clothing to prevent skin abrasions. Where necessary, use the extra padding provided with the device.

Ensure that the device does not put excessive pressure on bony prominences.

Skin integrity should be assessed before and after each use.

Use only input power cables, battery chargers, and battery assemblies offered by Ekso Bionics.

Ensure the EksoNR batteries are fully charged or adequate for the treatment session.

Appropriate spotting must be done at all times, as the EksoNR is not designed to prevent falling. The spotter must remain within an arm’s reach of the patient at all times while spotting. The role of the spotter is to provide guarding support to the patient while transferring, standing, walking, turning, and sitting. In an emergency, the spotter may help perform a long-legged sit maneuver.

New patients and patients at higher risk of falling require more supervision and spotting.

Always give clear, verbal instructions to patients, especially when beginning or changing an EksoNR action or when the patient is turning or is changing direction during EksoNR walking sessions.

Further research is needed to determine the long-term effects of walking with the EksoNR.

EksoNR session may cause fatigue and perhaps increase fall risk to certain users.

Be aware of leg length differences and/or bony abnormalities that would prevent safe operation. Use of the device should be paused and checked for fit if the patient complains of discomfort or pain. If the fit is correct and the patient continues to experience significant discomfort, this device may not be appropriate for use with this patient.

The EksoNR should be removed before entering a magnetic resonance (MR) scan room or having an MR scan. The EksoNR has not been evaluated for safety and compatibility in the MR environment. It has not been tested for heating, migration, or image artifact in the MR environment. The safety of EksoNR in the MR environment is unknown. Scanning a patient who has this device may result in patient injury. Contact with or being in proximity to an MR scanner may cause the exoskeleton to move or lead to electric shocks and may result in patient injury.

## Safety Reminders

### For Physical Therapists

An EksoNR-trained physical therapist must remain – or delegate another person who must remain—within arm’s reach of a patient at all times in case of a patient loss of balance or a device power loss.

EksoNR is NOT designed to prevent falling.

Always use the EksoNR with sufficient supervision and physical assistance.

New patients, and those at higher risk of falling, may require more supervision and physical assistance.

Always use the EksoNR with an assistive device.

Give clear verbal instructions and physical cueing to patients and your colleagues.

Provide clear instructions to patients whenever beginning or changing an EksoNR action and when stopping, turning, or changing direction. (For example, “beginning stepping now.”)

Provide clear instructions to your colleagues as you work in pairs. Verbalize that each partner is ready prior to initiating standing, walking, turning, or sitting during a patient session. Efficiently push the Stop button if your partner indicates a stop or pause is necessary.

Assess skin integrity before and after every training session. Special attention should be given to areas of EksoNR contact.

Patients should perform adequate lower extremity stretching before and after EksoNR sessions, as indicated by their physical therapist.

Discontinue use if the patient reports feeling discomfort or displays signs and symptoms of adverse reactions.

A dedicated spotter or tether is optional for additional safety with new or challenging patients. A safety suspension system may be attached to the EksoNR nylon cord but must be kept slack for normal EksoNR walking (no pulling, dragging, or suspension). Additionally, the suspension tether must be removed during sit-to-stand and stand-to-sit transitions.

A dedicated spotter is recommended when using FES interface.

### Device Specific Safety Considerations

Do not use the EksoNR if the device is not operating properly.

Discontinue use if the Controller directs.

Never use the EksoNR if the device has frayed or broken wires.

Never use the battery charger if the charger cord is frayed or compromised in any way.

Ensure the EksoNR batteries are fully charged or adequate for treatment session.

## Adverse Events

Any adverse reactions noticed by a physical therapist or patient during or following the use of the EksoNR should be documented thoroughly and reported to Ekso Bionics Customer Experience immediately. Examples of adverse reactions include:

Skin integrity issues

Bruising

Pain

Sudden nausea or changes in skin color, such as turning pale or red

Unusual swelling

Any undesired EksoNR mechanical behavior should be noted and reported to Ekso Bionics Customer Experience before continuing operation. Examples of undesired behaviors include:

Changes in normal walking or sit-to-stand/stand-to-sit transitions

Uneven step lengths related to software rather than patient performance

Unusual oscillations of any joint

Unusual sounds not present during training or regular EksoNR operation

## Training and Levels of Certification for EksoNR Use

The EksoNR is sent with GaitCoach software.

### GaitCoach Training

This is the training you will receive with the purchase of your EksoNR.

Introduce the feature set of the EksoNR.

Goal of training – Safe, effective, independent clinical use of the EksoNR

Certification goal - Level 2 certification for physical therapist(s)

- Can operate the EksoNR independently, or with any support personnel when additional support is needed.

  - If the therapist is unsafe to utilize device independently, they will be certified at Level 1, meaning that they can operate the EksoNR with anyone who has an equal or higher level of certification.

- May delegate EksoNR treatment session to support personnel at their discretion.

- When delegating tasks, the Level 2-trained physical therapist is still responsible for the patient, the EksoNR device, and for providing education to support personnel on device function and safety procedures.

- Physical therapist is able to provide training to other physical therapists on EksoNR with GaitCoach operation in preparation for EksoNR certification by Ekso Bionics

### GaitCoach+ Training

This training is available for purchase and will provide access to additional software features that will allow further customization of EksoNR therapy sessions. Trained physical therapists will maintain their level of certification after this training.

## Clinical Delegation

Ekso Bionics Requires the Following for Safe Delegation:

- A Level 1 or 2-trained physical therapist must provide the initial evaluation and initial training sessions for each patient.

- A Level 1 or 2-trained physical therapist is responsible for each EksoNR device session and maintain awareness of the patient’s status, condition, and EksoNR device settings while the device is in use.

- A Level 1or 2-trained physical therapist must provide the training necessary for the support staff to safely handle any patient.

Ekso Bionics acknowledges that a Level 2-trained physical therapist may choose to delegate the patient to support staff when deemed appropriate, but the physical therapist does this at their sole risk and responsibility. Ekso Bionics acknowledges that all physical therapists must exercise their own best clinical judgment in compliance with the laws and regulations governing their practice in determining when delegation may occur.

Safe delegation of EksoNR sessions requires appropriate training and patient selection for the patient and spotter to be safely managed during a session. Delegation without proper training may result in injury to the patient or staff member; therefore, Ekso Bionics recommends the following:

The patient that is delegated to support staff should:

- Be safely manageable during walking sessions in EksoNR.

- Be safely manageable with the assistive device and EksoNR programming.

- Be safely manageable with sit-to-stand and stand-to-sit transitions.

The delegate that is working with the patient should be:

- Trained sufficiently in EksoNR device operation to safely manage the patient in emergency operation procedures, such as Safe Mode, Pause, standing/sitting, etc.

- Able to manage the patient and the EksoNR device weight in the event of a loss of balance to prevent a fall.

Ekso Bionics utilizes a training checklist and competency template for the purposes of training certified physical therapists and strongly recommends the adoption of similar documentation for clinical delegates.

# Evaluation and Measuring for Fit

WARNING Importance of Evaluation

Each patient needs to be thoroughly evaluated, measured, and screened to determine appropriateness for EksoNR use. Physical therapists need to evaluate patient strength, range of motion, spasticity, skin integrity, and mobility. Failure to adequately evaluate a patient for appropriateness for EksoNR use could lead to patient injury, poor gait quality, and/or damage to the EksoNR. Accurate range of motion assessments are essential when determining appropriateness for EksoNR use.

# Turn EksoNR On and Off

Turn ON the EksoNR by pressing the Power Button on the back of the torso. The Button stays depressed when turned ON.

When the EksoNR is ready, the controller displays the Home screen.

The EksoNR may be turned OFF at any time by pressing the Power Button again; however, the patient should be removed from the EksoNR before turning OFF. When turned OFF, the hips go free and the knees lock in place.

NOTE: Ekso Bionics recommends that the EksoNR be turned off after each patient session to return the device to default settings and clear statistics.

# Errors

Patient training is essential for successful weight shifting and walking in EksoNR. Barriers to successful patient learning may be related to one or more of these areas.

- EksoNR adjustments:

- Lateral weight shift errors may be related to incorrectly set Leg Length or Hip Abduction.

Patient-specific issues:

- Spasticity or range of motion limitations in ankle may challenge forward weight shifting.

- Structural malalignment (for example, scoliosis) may make lateral weight shift more difficult.

- Vertical orientation errors or misinterpretation.

- Tense upper body movements may limit lateral and/or forward weight shifting.

- Patient using excessive compensatory methods to attempt to advance or clear lower extremities.

- Patient-EksoNR considerations:

- Loose backpack straps may result in suboptimal or challenging lateral or forward weight shifting.

- Excessive hip width of the device may cause inappropriate lateral weight shifting.

- Need for fit kit items for improved sagittal plane alignment within the EksoNR.

## Critical Device Errors

The EksoNR does the following when it detects an error:

Emits a beeping alert (from the top back portion of the EksoNR torso).

Displays an error code and message (often including instructions) on the controller.

### Safe Mode

WARNING Emergencies

In an emergency, press the Power button to release the hips and lock the knees, similar to Safe Mode. Then lower the patient and device to the ground using a long-legged sit.

Safe Mode is a safety feature that occurs if a critical device error happens. When the device enters Safe Mode, the EksoNR emits a distinct alarm cueing the physical therapists that immediate spotting is required for patient safety.

The following occurs during Safe Mode:

Hip motors go free with a damping component that slows the hip motion to assist the patient in staying upright (as long as EksoNR is still powered on), resisting the motion for hip flexion.

EksoNR knees lock in their current position.

Alarm sounds immediately and continuously until acknowledged on Controller.

Spotting patient involves:

Provide full extension moment through hips to provide mechanical stability.

    - Physical therapist stabilizes the patient’s hip by positioning physical therapist’s hip against the lower torso of the device.

    - Hold Torso Handle to extend torso OR

    - Place one hand on Torso Handle and one hand on Sacral Handle.

      - Pull back on Torso Handle and push forward on Sacral Handle to position patient’s hips anterior to shoulders.

      - Second physical therapist spots as necessary.

    - When safe to do so, read the message on the controller and note the error code.

Press the  button to stop the audible alarm (beeping alert) and to acknowledge the critical device error.

Hold the  button to clear the message and error code number from the screen.

Correct the error as instructed.

If unable to clear message, begin long-legged sitting:

- Emphasize communication among team (physical therapists and patient).

- Lead physical therapist provides primary support, standing behind patient.

- Second person assists patient as necessary.

- Lead physical therapist instructs second person to retrieve chair when patient is in a safe position.

- Chair positioned behind patient back to estimated length of lower extremities.

- Lead physical therapist stands to side while second person assists with sitting.

- Lead physical therapist controls lowering of patient into the chair while legs stay locked in their current position.

- Position patient upright in the chair; make sure patient is seated squarely in the chair.

- Second person moves to support lower extremities (see photo).

- Lead physical therapist uses Controller screen to navigate into Legs Free mode when safe.

- Position patient’s feet squarely on floor to stabilize in sitting.

If the error message reoccurs after making the appropriate adjustments, discontinue EksoNR use and contact Ekso Bionics Customer Experience.

EksoNR is in Safe Mode when:

Immediately powering on the device, before selecting Legs Free from walk menu.

When device is powered OFF.

When an alarm is signaled, that results in immediate Safe Mode.

NOTE: If the physical therapist wishes to purposely enter Safe Mode to lower a patient directly to the ground—in the event of severe medical event requiring immediate attention in supine—this is done through turning OFF the EksoNR through the Power button. In this situation, no damping is applied to the hip motors and they are fully free.

General considerations:

Hips go free in the event that the physical therapist needs to lower the patient all the way down to the floor.

- In the event of a device issue, the physical therapist needs to overcome the dampening motor resistance while lowering the patient using a long-legged sit.

- Depending on the patient’s presentation, the patient’s ability to manage Safe Mode varies:

    - Patient with trunk and lower extremity weakness or paralysis require full assistance to stabilize in Safe Mode.

    - Physical therapist needs to provide stability through the EksoNR hips, placing the EksoNR in an extension moment through hips, maintaining mechanical stability.

    - Patient with lower extremity strength with ability to ambulate over ground may be able to contribute some amount of stability upright when the EksoNR is in Safe Mode.

    - Physical therapist should still provide assistance to maintain the EksoNR torso in an upright position.

# Doffing

The physical therapist needs to help the patient doff the EksoNR after each session is concluded. The patient must be sitting in a chair or on a mat table that can support both the EksoNR and the patient during the entire doffing procedure. The EksoNR does not need to be turned OFF for doffing.

EksoNR straps can be loosened and unfastened in any order; however, maintaining patient balance and stability is crucial.

Make sure that the patient does not hit their buttocks on the hip abduction joint or the thigh support brace during the transfer.

Physical therapist positioned in front of the patient.

Start at the top and work down to the Foot Binding straps:

- Release the Chest strap.

- Loosen the Backpack straps and the Torso Pad.

- Unfasten the Thigh and Lower Leg straps.

- Release the Foot Binding straps.

Release the EksoNR legs into abduction by pressing the silver hip abduction tab.

- May require slight adduction pressure to release.

Remove the patient’s legs from the device.

Internally rotate EksoNR leg to allow for a clean path to transfer.

Document statistics of the session.

Power OFF device at the end of each patient session.

- Powering OFF clears all statistics and returns settings to default.

# Regulatory Information

## Medical device directive

0044

This product conforms to the requirements of council directive 93/42/EEC concerning medical devices when it bears the following CE mark of conformity:

# Capacities

| Capacities |  |
| --- | --- |
| User Weight | Maximum 220 lb (100 kg) |
| User Height | Approximately 5 ft, 0 in. to 6 ft, 4 in. (1.53 m to 1.95 m) |
| User Hip Width | Maximum 18 in. (45.6 cm) |
| Performance |  |
| Expected Service Life | 4 years or 1.25 million steps |
| Walking |  |
| Step Length | 8 to 18 in. (20 – 46 cm) |
| Step Height | Approximately 0.0 to 1.5 in. (0 to 3.8 cm) |
| Swing Time | 0.8 to 2.0 seconds |
| Joint Movements |  |
| Hip | Range of motion: 10º extension to 135º flexion Abduction: -2º to 4º |
| Knee | Range of motion: 0º extension to 130º flexion |
| Ankle | Range of motion: 10º dorsiflexion to 10º plantarflexion Resting Angle: 3º to 8º Dorsiflexion stiffness: Settings 1 to 4 |
| Batteries |  |
| Materials | Li-Ion |
| Voltage | 48.1 VDC |
| Capacity | 2000 mAh |
| Charging Time | Approximately 60 minutes |
| Power Source | 110 - 220 VAC, 50 - 60 Hz |
| Safety | IEC 62133 |
| Charger Specifications |  |
| Input power | 110 - 220 VAC, 50 - 60 Hz |
| Safety | IEC 60601-1:2005 |
| EMC standards | IEC 60601-1-1:2007 |
| Environmental |  |
| Operating temperature range | 41ºF to 104ºF (5ºC to 40ºC) |
| Storage, Shipping, and Transport temperature range | -12ºF to 158ºF (-25ºC to 70ºC) NOTE: If storage or transport temperature is outside the operating temperature range, allow a minimum of 4 hours for device temperature to stabilize in the operating environment. |
| Relative Humidity range | 15 to 90% RH |
| Atmospheric Pressure range | 700 to 1060 hPa |
| Operating temperature range | 41ºF to 104ºF (5ºC to 40ºC) |

WARNING Communication

In a medical facility, this product can only be used in areas permitting wireless or cellular communication.

# Batteries

The EksoNR uses two 48-volt rechargeable batteries. The batteries attach to each side of the EksoNR torso when pins and other positioning pieces are in proper alignment. EksoNR batteries work together as a set of two. Each pair should be replaced and/or recharged simultaneously to maximize battery life and ensure proper functionality of the EksoNR.

When batteries are attached to the EksoNR torso and the device is turned ON, the controller displays the power level in the upper right corner.

Check the controller to determine if the batteries are sufficiently charged prior to each patient session.

Check that both batteries are properly attached to the torso prior to each patient session.

Always use two batteries during EksoNR operations

Respond promptly to Low Battery warnings displayed by the controller.

Remove or replace batteries only when the EksoNR is turned OFF and in a sitting position.

Recharge each battery at least once every two months.

Use only the two Ekso Bionics-provided Battery Chargers to charge batteries.

Always check that a Battery Charger is plugged in before plugging in a battery for recharging.

Never use the Battery Chargers if the charger cord is frayed or compromised in any way.

CE NOTICE - INFORMATION FOR THE USER

This equipment has been tested and found to comply with the limits for a Class B digital device, pursuant to Part 15 of the FCC rules. This device is also found to comply with the limits for a Class B digital device, pursuant to presently harmonized European standards. These limits are designed to provide reasonable protection against harmful interference when the equipment is operated in a commercial environment. This equipment generates, uses, and can radiate radio frequency energy and, if not installed and used in accordance with the instruction manual, may cause harmful interference to radio communications. Operation of this equipment in a residential area is likely to cause harmful interference, in which case the user will be required to correct the interference at the expense of the user.

The user may find the following publication prepared by the Federal Communications Commission helpful:

How to Identify and Resolve Radio-TV Interference Problems (Stock Number 004-000-00345-4), available exclusively from the Superintendent of Documents, Government Printing Office, Washington, DC 20402 (telephone 202-512-1800).

WARNING

Changes or modifications not expressly approved by the party responsible for compliance to Part 15 of the FCC Rules could void the user’s authority to operate the equipment.

This device complies with Part 15 of the FCC Rules. Operation is subject to the following two conditions:

- This device may not cause harmful interference.

- This device must accept any interference received, including interference that may cause undesired operation.

Caution Legal

- Federal law restricts this device to sale by or on the order of a physician.

- No operator-serviceable parts inside unit.

- Refer servicing to qualified personnel.

- Use only data cables offered by Ekso Bionics, Inc. for this system.

EksoNR Unit Precautions and Warnings

WARNING Compliance

Changes or modifications not expressly approved by the party responsible for compliance to Part 15 of the FCC Rules could void the user’s authority to operate the equipment.

EksoNR Battery Precautions and Warnings

Caution Batteries

- Do not use or leave the battery in a place exposed to direct sunlight or in a car unattended.

- Follow the instructions given in this manual regarding battery installation.

- Fully charge the battery before using the first time. Do this by carefully reading and following the instructions specified for the charger and learn how to charge the battery correctly.

- To save battery life, turn off your equipment power switch after each use.

- Replace the battery if alerted by a low battery warning.

WARNING Batteries

- Use the battery in only the following conditions:

  - In Use: -4 to +140ºF (-20 to +60ºC)

  - While Charging: 32 to 113º F (0 to 45ºC)

  - Storage: The battery may be stored up to 2 months at room temperature—68ºF ± 5º (20ºC +/- 5º)—without charging.

- Do not put the battery into a fire or heat it.

- Do not connect any battery terminals with metal objects, such as wire.

- Keep out of reach of children.

- Do not allow the battery to get wet.

- Do not throw, drop, hammer, stamp, or impact the battery.

- Do not disassemble or alter the battery.

- Do not use the battery if it appears damaged.

# Guidance and Manufacturer’s Declaration

Emissions

All Medical Electrical Equipment and Medical Electrical Systems

Guidance and Manufacturer’s Declaration – Electromagnetic Emissions

The EksoNR is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment

| Emissions Test | Compliance | Electromagnetic Environment – Guidance |
| --- | --- | --- |
| RF Emissions CISPR 11 | Group 1 | The EksoNR uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF Emissions CISPR 11 | Class A | The EksoNR is suitable for use in all establishments other than domestic and those directly connected to the public low-voltage power supply network that supplies buildings used for domestic purposes |
| Harmonics Emissions EN 61000-3-2 | Not Applicable1 |  |
| Voltage Fluctuations/ Flicker Emissions EN 61000-3-3 | Not Applicable1 |  |

1The EksoNR is battery powered and does not contain any AC or DC power ports. The EksoNR does not contain any input/output lines longer than 3 m.

Immunity

All Medical Electrical Equipment and Medical Electrical Systems

Guidance and Manufacturer’s Declaration – Electromagnetic Immunity

The EksoNR is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment

| Immunity Test | IEC 60601Test Level | Compliance Level | Electromagnetic Environment – Guidance |
| --- | --- | --- | --- |
| ESD IEC 61000-4-2 | ± 8 kV Contact ± 15 kV Air | ± 8 kV Contact ± 15 kV Air | Floors should be wood, concrete, or ceramic tile. If floors are synthetic, the r/h should be at least 30%. |
| EFT IEC 61000-4-4 | ± 2 kV Mains ± 1 kV I/O lines | Not Applicable1 | The EksoNR is battery powered. |
| Surge IEC 61000-4-5 | ± 1 kV Differential ± 2 kV Common | Not Applicable1 | The EksoNR is battery powered. |
| Voltage Dips/Dropout IEC 61000-4-11 | >95% Dip for 0.5 Cycle, >95% Dip for 1 Cycle, 30% Dip for 25/30 Cycles, >95% Interruption for 250/300 Cycles | Not Applicable1 | The EksoNR is battery powered. |
| Power Frequency 50/60 Hz Magnetic Field IEC 61000-4-8 | 30 A/m | 30 A/m | Power frequency magnetic fields should be that of a typical commercial or hospital environment. |

1The EksoNR is battery powered and does not contain any AC or DC power ports. The EksoNR does not contain any input/output lines longer than 3 m.

NOTE: The EMISSIONS characteristics of this equipment make it suitable for use in industrial areas and hospitals (CISPR 11 class A). If it is used in a residential environment (for which CISPR 11 class B is normally required) this equipment might not offer adequate protection to radio-frequency communication services. The user might need to take mitigation measures, such as relocating or re-orienting the equipment.

Immunity

Medical Electrical Equipment and Medical Electrical Systems that are NOT Life-supporting

Guidance and Manufacturer’s Declaration – Electromagnetic Immunity

The EksoNR is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment.

| Immunity Test | IEC 60601 Test Level | Compliance Level | Electromagnetic Environment – Guidance |
| --- | --- | --- | --- |
| Conducted RF IEC 61000-4-6 | 3 Vrms 150 kHz to 80 MHz | Not Applicable1 | Portable and mobile communications equipment should be separated from the EksoNR by no less than the distances calculated/listed below: |
|  |  |  | d = 1.2 ÖP 150 kHz to 80 MHz |
| Radiated RF IEC 61000-4-3 | 3 V/m 80 MHz to 2.7 GHz | 3 V/m | d = 1.2 ÖP 80 to 800 MHz d = 2.3 ÖP 800 MHz to 2.7 GHz |
|  |  |  | where P is the maximum output power rating of the transmitter in watts (W), and d is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, as determined by an electromagnetic site surveya, should be less than the compliance level in each frequency rangeb. Interference may occur in the vicinity of equipment containing a transmitter. |
| NOTE 1: At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects, and people. |  |  |  |
| 1The EksoNR is battery powered and does not contain any AC or DC power ports. The EksoNR does not contain any input/output lines longer than 3 m. a Field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast, and TV broadcast, cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the EksoNR is used exceeds the applicable RF compliance level above, the EksoNR should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as reorienting or relocating the EksoNR. b Over the frequency range 150 kHz to 80 MHz, field strengths should be less than 3 V/m. |  |  |  |

Recommended Separation Distances Between Portable and Mobile RF Communications Equipment and the EksoNR

Medical Electrical Equipment and Medical Electrical Systems that are NOT Life-supporting

Recommended Separations Distances for the EksoNR

The EksoNR is intended for use in the electromagnetic environment in which radiated disturbances are controlled. The customer or user of the EksoNR can help prevent electromagnetic interference by maintaining a minimum distance between portable and mobile RF Communications Equipment and the EksoNR as recommended below, according to the maximum output power of the communications equipment.

| Max Output Power (Watts) | Separation (m) 150 kHz to 80 MHz d = 1.2 ÖP | Separation (m) 80 to 800 MHz d = 1.2 ÖP | Separation (m) 800 MHz to 2.7 GHz d = 2.3 ÖP |
| --- | --- | --- | --- |
| 0.01 | 0.12 | 0.12 | 0.23 |
| 0.1 | 0.38 | 0.38 | 0.73 |
| 1 | 1.2 | 1.2 | 2.3 |
| 10 | 3.8 | 3.8 | 7.3 |
| 100 | 12 | 12 | 23 |

For transmitters rated at a maximum output power not listed above, the recommended separation distance d in meters (m) can be estimated using the equation applicable to the frequency of the transmitter, where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer.

NOTE 1: At 80 MHz and 800 MHz, the separation distance for the higher frequency range applies.

NOTE 2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects, and people.

Emissions

All Medical Electrical Equipment and Medical Electrical Systems

Guidance and Manufacturer’s Declaration – Electromagnetic Emissions

The EksoNR Battery Charger is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment

| Emissions Test | Compliance | Electromagnetic Environment – Guidance |
| --- | --- | --- |
| RF Emissions CISPR 11 | Group 1 | The EksoNR Battery Charger uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF Emissions CISPR 11 | Class B | The EksoNR Battery Charger is suitable for use in all establishments, including domestic establishments and those directly connected to the public low-voltage power supply network that supplies buildings used for domestic purposes. |
| Harmonics EN 61000-3-2 | Class A |  |
| Flicker EN 61000-3-3 | Complies |  |

Immunity

All Medical Electrical Equipment and Medical Electrical Systems

Guidance and Manufacturer’s Declaration – Electromagnetic Immunity

The EksoNR Battery Charger is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment.

| Immunity Test | IEC 60601 Test Level | Compliance Level | Electromagnetic Environment – Guidance |
| --- | --- | --- | --- |
| ESD IEC 61000-4-2 | ± 8 kV Contact ± 15 kV Air | ± 8 kV Contact ± 15 kV Air | Floors should be wood, concrete, or ceramic tile. If floors are synthetic, the r/h should be at least 30%. |
| EFT IEC 61000-4-4 | ± 2 kV Mains ± 1 kV I/O lines | ± 2 kV Mains ± 1 kV I/O lines | Mains power quality should be that of a typical commercial or hospital environment. |
| Surge IEC 61000-4-5 | ± 1 kV Differential ± 2 kV Common | ± 1 kV Differential ± 2 kV Common | Mains power quality should be that of a typical commercial or hospital environment. |
| Voltage Dips/Dropout IEC 61000-4-11 | >95% Dip for 0.5 Cycle, 60% Dip for 5 Cycles, 30% Dip for 25 Cycles, >95% Dip for 5 Seconds | >95% Dip for 0.5 Cycle, 60% Dip for 5 Cycles, 30% Dip for 25 Cycles, >95% Dip for 5 Seconds | Mains power quality should be that of a typical commercial or hospital environment. If the user of the EksoNR Battery Charger requires continued operation during power mains interruptions, it is recommended that the EksoNR Battery Charger be powered from an uninterruptible power supply or battery. |
| Power Frequency 50/60 Hz Magnetic Field IEC 61000-4-8 | 3 A/m | 3 A/m | Power frequency magnetic fields should be that of a typical commercial or hospital environment. |

Immunity

Medical Electrical Equipment and Medical Electrical Systems that are NOT Life-supporting

Guidance and Manufacturer’s Declaration – Electromagnetic Immunity

The EksoNR Battery Charger is intended for use in the electromagnetic environment specified below. The customer or user of the EksoNR should ensure that it is used in such an environment.

| Immunity Test | IEC 60601 Test Level | Compliance Level | Electromagnetic Environment – Guidance |
| --- | --- | --- | --- |
| Conducted RF IEC 61000-4-6 | 3 Vrms 150 kHz to 80 MHz | 3 Vrms | Portable and mobile communications equipment should be separated from the EksoNR Battery Charger by no less than the distances calculated/listed below: |
|  |  |  | d = 1.2 ÖP 150 kHz to 80 MHz |
| Radiated RF IEC 61000-4-3 | 3 V/m 80 MHz to 2.5 GHz | 3 V/m | d = 1.2 ÖP 80 to 800 MHz d = 2.3 ÖP 800 MHz to 2.5 GHz |
|  |  |  | Where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer, and d is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, as determined by an electromagnetic site surveya, should be less than the compliance level in each frequency rangeb. Interference may occur in the vicinity of equipment containing a transmitter. |
| NOTE 1: At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects, and people. |  |  |  |
| a Field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast, and TV broadcast, cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the EksoNR Battery Charger is used exceeds the applicable RF compliance level above, the EksoNR Battery Charger should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as reorienting or relocating the EksoNR Battery Charger. b Over the frequency range 150 kHz to 80 MHz, field strengths should be less than 3 V/m. |  |  |  |

Recommended Separation Distances Between

Portable and Mobile RF Communications Equipment and the EksoNR Battery Charger

Medical Electrical Equipment and Medical Electrical Systems that are NOT Life-supporting

Recommended Separations Distances for the EksoNR Battery Charger

The EksoNR Battey Charger is intended for use in the electromagnetic environment in which radiated disturbances are controlled. The customer or user of the EksoNR can help prevent electromagnetic interference by maintaining a minimum distance between portable and mobile RF Communications Equipment and the EksoNR as recommended below, according to the maximum output power of the communications equipment.

| Max Output Power (Watts) | Separation (m) 150 kHz to 80 MHz d = 1.2 ÖP | Separation (m) 80 to 800 MHz d = 1.2 ÖP | Separation (m) 800 MHz to 2.7 GHz d = 2.3 ÖP |
| --- | --- | --- | --- |
| 0.01 | 0.12 | 0.12 | 0.23 |
| 0.1 | 0.37 | 0.37 | 0.73 |
| 1 | 1.2 | 1.2 | 2.3 |
| 10 | 3.8 | 3.8 | 7.3 |
| 100 | 12 | 12 | 23 |